## Slide 1: 3d-Printed Granular Jamming Hand
Melissa Indoe
Chris Kang
Sean Phelan
Maggie Serra
Chris Wallace

## Slide 2: Problem Statement
There are 3 million arm amputees worldwide
  59% have amputations below the elbow
Current prosthetic hands are costly
  Only half of arm amputees can afford prosthetic hands
  People in developing countries cannot get access to such devices
There is a global need to develop affordable, readily available prosthetics

## Slide 3: Project Objective
Create a “below-the-elbow” prosthetic hand
The hand is:
  Open Source
  Highly Affordable (Costing $363)
  High-Functioning
  Task-Oriented
  Simple to Produce

## Slide 4: Deliverables
Completed Prototype
  Hand
  Granular Jamming Pad
  Forearm
    Electronics
    Arduino Program
Website
  SquareSpace: Project Information
  Github: Downloadable Source Files
Team Poster

## Slide 5: Device Design Overview
Hand Design
Forearm Design
Servo Hub Design
Granular Jamming Pad
Electronics and Programming

## Slide 6: 1. Hand Design
Matches size and weight of human hand as closely as possible
Thumb is mounted in opposition to middle finger
Actuated by cables
Fingers are held straight by dental rubber bands
  First joint of each finger uses two rubber bands
  All other joints use one rubber band

## Slide 7: 1. Hand Design
Design Updates:
Redesigned cable guides
  U-shape for large contact area
  Holes for support material removal
  Updated routing for thumb
Hole for granular jamming fitting
  Hexagonal to prevent rotation
Set screw support
  Prevents hyperflexion
  Creates more fluid motion
Old (left) and new (right) designs.
Bottom: Added set screw support.

## Slide 8: 2. Forearm Design
Design Updates:
Motor mounts updated
Sensor mounts added
Hand mount has been moved to lower half of forearm
Forearm has been separated into segments for ease of printing
Arduino mount has been reduced in size
Top: Forearm design at the conclusion of phase 5.
Bottom: Final forearm design.

## Slide 9: 2. Forearm Design
Sensor and motor mount:
Hole for fishing line
  Keeps fishing line level with potentiometer
Potentiometer mount
  Insets for pins on both sides
    Allows for forward/backward mounting
  Depressions to protect solder points
Motor mount
  Holes for wires on both sides
    Allows for forward/backward mounting
Top: Sensor and motor mount model.
Bottom: Two assembled sensor and motor mounts.

## Slide 10: 2. Forearm Design
Bottom Assembly:
Attachment points for:
  Servo motors
    Servo hubs
  Potentiometers
  Hand
Printed in two pieces
  Accommodates print bed
    Arm is 12 inches long
    Print space available is 10x10x10 inches
Bottom forearm half segments (top) and fully assembled half forearm (bottom).

## Slide 11: 2. Forearm Design
Top assembly:
Attachment point for Arduino
Encloses forearm
  Bolted to lower half of forearm at four points
  Also could be sealed to lower half with ABS cement
Printed in two pieces
Top forearm half segments (top) and fully assembled half forearm (bottom).

## Slide 12: 3. Servo Hub Design
Smallest diameter possible
  Minimum diameter was calculated for desired range of motion
Staggered heights
  Able to fit over and under each other within forearm housing
Fishing line tunnel
  Fishing line is fed into tunnel and tied off at top of hub
  This keeps line from sliding off flanges
Wireframe model of servo hubs (top) and printed hubs in use (bottom)

## Slide 13: 3. Servo Hub Design
Design Challenge:
  3D printing teeth to fit the servo gear is almost impossible
    Servos come with injection-molded hubs (higher tolerance)
  Staggered heights prevents use of included hubs
  Solution would have to be simple enough for a layperson
Solution:
  The solution was a 5 step process taking just a few minutes:
    3D print hubs with holes slightly smaller than servo gears
    Use ABS cement to soften plastic
    Press hubs onto servo gears
    Pull hubs off gears
    Wait one minute for plastic to re-solidify

## Slide 14: 4. Granular Jamming Pad
Previous Design:
  Reusable latex kitchen gloves secured into a palm shape using super glue
Design issues:
  Super glue cracked, damaging the airtight seal
Granular jamming pad demonstrating cracking.

## Slide 15: 4. Granular Jamming Pad
New Design:
  Reusable latex kitchen gloves secured into a palm shape using Loctite brand superflex silicone adhesive sealant
Pros:
  Creates a better seal than super glue
  Seal is flexible and will not crack
Cons:
  Takes 24 hours to fully set
  Creates thick, somewhat inflexible seams
Granular jamming pad sealed with silicone adhesive.

## Slide 16: 4. Granular Jamming Pad
  Incorporation into the hand:

## Slide 17: 5. Electronics and Programming

## Slide 18: 5. Electronics and Programming

## Slide 19: 5. Electronics and Programming
Program has changed
  No longer MATLAB based
  Arduino only
Control of vacuum with push button
Control of servos with rotary potentiometers
Linear potentiometers add capability to contour around slanted surfaces
  Servos Move  check if potentiometers have updated
  If Yes  move next step
  If No  do not move

## Slide 20: Prototype Fabrication
Fabrication steps are focused on simplicity
  Assembly requires only screwdrivers and ABS cement
  Hand is non-assembly, forearm is assembled in eight parts
Assembly instructions are provided on github
  Instructions for pad, forearm, and hand assembly are all available
  Electronics schematic is also available

## Slide 21: Prototype Fabrication
Fabrication Steps:
Motors and sensors are secured into motor mounts
Motor mounts are dropped into forearm and bolted in
Hand is cabled and rubber bands are added
Hand is screwed to forearm
Granular jamming pad is secured to hand
Arduino is bolted to top half of forearm
Cables are run to motor hubs
Wiring is hooked up according to schematic
Forearm is bolted together

## Slide 22: Problems in Fabrication
Potentiometers added friction
  Solution: Extra rubber bands at finger bases
  Mounts did not line up with servos
    Fixed in the model but not the prototype
  Printer did not print hand mount due to malfunction
    Printed as a standalone hand mount and secured to arm via ABS cement
  In general, very few assembly issues
    This shows that the assembly is very straightforward and simple

## Slide 23: Performance Testing
Finger range test
  Actually limited slightly by length of potentiometers
  Potentiometers also created additional friction
    Sometimes caused the fingers to stay slightly curled when resting
Demo control
  Showed functional contact detection
  Established control of finger position
Granular jamming pad test
  Showed that designed pad was slightly too stiff and shallow
  Balloon latex proved to be a more effective material

## Slide 24: Performance Testing: Finger Range & Demo Control

## Slide 25: Performance Testing: Granular Jamming Pad

## Slide 26: Additional Performance Testing
Unfortunately, due to an electrical short, this is where performance testing concluded
Additional tests that would have been conducted:
  Finger strength
    Close hand around objects of increasing weight, attempt to lift objects off table without fingers slipping
  Finger dexterity
    Close fingers around a non-uniformly shaped object
    Success will be measured by how natural the grip appears, and how difficult it is to break grip and pull the object out of the hand
  Potentiometer sensitivity
    Close hand around a fragile object (such as a paper cup)
    If fingers can close without crushing, test will be successful

## Slide 27: Potential Improvements
Replace demo code with Bluetooth control
Additional updates to the forearm
  Design in an open source program to make files easier to edit
  Test print on a low-end printer without support material
  Fit control elements more efficiently
Consider attachment to a person
  Design 3D printed socket with customizable fit
Update granular jamming pad
  Develop a pad with more depth
  Utilize a more flexible housing material

## Slide 28: Gantt Chart

## Slide 29: Budget
Budget as spent by the senior design team:

## Slide 30: Budget
Budget for manufacture by an end user:

## Slide 31: Website
Granular3dhand.squarespace.com
  Project and team information
  Provides blog-style project updates
Github.com/3D-Printed-Granular-Jamming-Hand
  Provides source code and parts files
  Full history of all file updates and changes